1. Introduction

Leak testing of sealed sources is mandated under CNSC regulations to ensure that sealed sources have not developed defects, sustained damage, or degraded in a manner that could present an unrecognized radiological risk to individuals using or working near them. This SOP aligns with the CNSC's REGDOC-1.6.1, Appendix AA, and Queen's University's Radiation Safety Manual.

2. Scope

This procedure applies to all laboratories at Queen’s University that possess sealed sources containing 50 MBq or more of a nuclear substance or nuclear substances used as shielding.

A centralized spreadsheet maintains an up-to-date record of all sealed sources at the university requiring leak testing. It includes:

- Location of each sealed source

- Required leak testing frequency for each source

A secondary tab provides calculated Minimum Detectable Activities (MDAs) for the specific radioisotopes measured during leak testing.

Spreadsheet Access Path:  \\N:\EHSShare\Radiation\1 - Ionizing Radiation\1 - Canadian Nuclear Safety Commission\3 - Regulatory Files\3 - Inventory\Sealed Sources & Devices\2 - Leak Testing

3. Applicable Legislation, Standards, Guidelines

This procedure complies with the following regulatory requirements for leak testing of sealed sources:

- Nuclear Substances and Radiation Devices Regulations, Section 18

- REGDOC-1.6.1, Appendix AA: Regulatory Expectations for Leak Testing of Sealed Sources

4. Responsibilities

Radiation Safety Officer (RSO):

- Administer, control, and update this procedure and associated documentation.

- Maintain the university-wide sealed source inventory.

- Ensure timely leak testing as per spreadsheet schedule.

- Review and approve Sealed Radiation Source Leak Testing Certificates.

- Provide training to authorized personnel.

- Investigate results exceeding action limits.

- Notify CNSC of any reportable conditions under SOP-Radiation 04.

Permit Holder:

- Ensure appropriate facilities, equipment, instruments, supervision, and instruction.

- Report all radiation incidents to the RSO per Radiation Committee Policy.

Designated Employee:

- Complete University Radiation Safety Training before working in labs.

- Be familiar with Section 18 of the Nuclear Substances and Radiation Devices Regulations.

- Only perform leak tests after specific training and authorization by the RSO.

- Complete Annual Radiation Refresher Training.

- Report any of the following to their supervisor:

  - Increased environmental or health risks

  - Security threats or breaches

  - Regulatory non-compliance

  - Sabotage, theft, loss, or misuse of nuclear substances

  - Releases exceeding permit levels

5. Definitions

- Sealed Source: A radioactive material encapsulated in a solid form to prevent release. Used for calibration, irradiation, or measurement.

- Leakage Test: A method of confirming the integrity of a sealed radioactive capsule by detecting any potential radioactive contamination.

6. Safety Precautions

- Treat the device as contaminated during the wipe.

- Use two layers of nitrile gloves and long forceps (recommended).

- Treat all wipes as contaminated until verified clean (<200 Bq).

- Avoid cross-contamination.

- Wash hands thoroughly after testing.

7. Work Practices

Leak testing will be conducted according to manufacturer-specific procedures outlined in Appendix A. Testing involves two main steps:

- Wiping the source

- Measuring the wipe sample

7.1 Wipe Testing Materials

- Wipe the sealed source device (where leakage could occur) with a dry cotton swab.

- Use Whatman Glass Microfibre Filters (dissolves in scintillation fluid).

- Use 20 mL plastic scintillation vials, labeled with source identifiers.

- Insert filters into vials, add ~15 mL scintillation cocktail, cap securely, and shake.

- Use a calibrated survey meter to scan the source after wiping.

7.2 Measurement Equipment

- Use a Beckman LS6500 Liquid Scintillation Counter (Serial: 7069021), located in Botterell Hall, Room 532.

- Select User Program with open window (e.g., Channel 1000) for the radionuclide.

- Run calibration before use—refer to LSC Calibration Procedure.

- Refer to Beckman documentation for isotope spectra and counting efficiencies.

- Report wipe test results to the RSO and Permit Holder.

- Maintain all records at the EH&S office.

7.3 Analysis of Results

- Results exceeding 150 CPM trigger an RSO-led investigation.

- Results above 200 Bq are reported to the CNSC immediately.

8. Training Requirements

- Training for sealed source leak testing will be provided by the RSO.

9. References

- CNSC (2011). Regulatory Expectations for Leak Testing of Sealed Sources, February 7.

10. Revision History

| Date | Revised by | Changes |
| --- | --- | --- |
| 05/21/2025 | Raico Lamela | Format and Sections content updated |

Appendix A

Queen’s University

Environmental Health & Safety

Sealed Source Leak Testing Procedures

Sealed Source Wipe Testing Procedure:

| Permit # |  | Permit Holder |  |
| --- | --- | --- | --- |
| Building |  | Room |  |
| Manufacturer |  |  |  |
| Serial # |  | Model # |  |
| Isotope |  | Activity |  |

Safety Precautions (Use of PPE, badger or any other device and conditions to reduce exposure to radiation leakage from the source ALARA at the time of testing the source )

1.

2.

3.

4.

Leak test instructions (must be written following the manufacturer manual)

1.

2.

3.

4.

# Appendix 13 – Sealed Radiation Source Leak Testing Certificate

Queen’s University	Contact Person: Jamie Coad, RSO Environmental Health & Safety	Phone: (613) 533-6000 78358          355 King Street West, 1st Floor, West Wing	Phone 24 Hours: (613) 533-6111

Kingston, ON, K7L 3N6	(Queen’s Emergency Report Centre) CNSC Licence:

Wipe Sampler’s Information

Name:  	 Address:

Phone:  	 Signature:

Source Information

Permit #:	Permit Holder:   	 Building:	Room:   	 Manufacturer:  		 Serial #:	Model #:   	 Isotope:	Activity:

| VIAL # | CPM | BKGR. | EFF. FACTOR | DPM | <200 Bq | >200 Bq |
| --- | --- | --- | --- | --- | --- | --- |

Wipe Measurement Information

Measured by:			Date:    	                        Phone: ____________     Signature:   	______________________________

Action to be Taken

Under 150 dpm - results to file

>150 dpm and <200 Bq: Investigation by RSO

>200 Bq - isolate source and inform CNSC.